## Sheet1
Clause | Requirement Summary | Applicable? (Y/N) | Current Procedure/Doc | Objective Evidence (Records to sample) | Status (C/PC/NC) | Nonconformity Description | Risk Impact (H/M/L) | Markets (US/EU/IR) | QMSR-specific? (Y/N) | FDA-Visible Record? (MR/Internal/Supplier) | Owner | Due Date | Remediation Action | Acceptance Criteria (Verification) | CAPA ID | Verification Method | Sample Plan
4.1.5 (Outsourced processes) | Control outsourced processes to ensure conformity; maintain QMS control over third parties. | Y | P-740 Supplier control; Q-AGMT-STER-01 Quality Agreement (Sterilizer) | Approved supplier list, qualification pack, latest audit report, COC/COA, deviations/CAPA from suppliers | PC | Supplier re-evaluation periodicity not defined for high-risk processes; no KPI trend for sterilizer OOS. | H | US, EU, IR | Y | Supplier | Supplier Quality Manager | 2025-10-15 | Define risk-based re-evaluation frequency; add OTD/OOS KPIs; update quality agreement; plan supplier audit. | Revised P-740 v2.0; KPI dashboard live; executed audit with no major NCs. | CAPA-2025-014 | Doc review + record sampling + interview | All critical suppliers + 3 non-critical at random
5.6 (Management review) | Top management reviews QMS at planned intervals; inputs/outputs documented. | Y | P-560 Management Review; MR-Minutes-2024-12 | MR agendas, minutes, actions, KPI packs | PC | Not all required inputs (supplier performance, PMS trends) included; action tracker lacks due dates. | M | US, EU, IR | Y | MR | MR / QA Head | 2025-11-01 | Add missing inputs; implement action tracker with owners & deadlines; schedule interim MR before QMSR go-live. | MR pack checklist complete; minutes show actions with owners/dates; closure evidence. | CAPA-2025-019 | Doc review + interview | Last 2 MRs
6.2 (Human resources) | Personnel competence; training; awareness of regulatory requirements. | Y | P-620 Training & Competence; Competency Matrix v1 | Training records, effectiveness checks, job descriptions | C |  | M | US, EU, IR | N |  | HR / QA | — | — | — |  | Record sampling | 3 per department
7.3 (Design & development) | Plan, inputs, outputs, reviews, verification, validation, transfer, changes. | Y | P-730 Design control; D&D plan templates; RMF per ISO 14971 | Design plan, DI/DO traceability, V&V reports, usability (IEC 62366), software lifecycle (IEC 62304) if applicable | PC | Usability formative evidence missing for Class II device; software SOUP assessment incomplete. | H | US, EU | N |  | R&D Manager | 2025-12-01 | Plan and execute formative usability study; finalize SOUP assessments; update trace matrix. | Approved reports linked to DMR/MDF; trace matrix complete. | CAPA-2025-023 | Doc review + interview | Last 2 projects
7.5.6 (Validation of processes) | Validate special processes (where output cannot be fully verified). | Y | P-756 Process Validation; IQ/OQ/PQ templates | Validation master plan, IQ/OQ/PQ for sealing/sterilization, requalification schedule | NC | No requalification for heat sealer since 2023; PQ lacks worst-case configuration. | H | US, EU, IR | N | Internal | Production Eng. Lead | 2025-10-05 | Execute overdue requalification; update PQ including worst-case; implement annual schedule. | Signed IQ/OQ/PQ; schedule approved; first periodic check done. | CAPA-2025-027 | Record sampling + observation | All special processes
7.5 (Packaging / ISO 11607) | Validated sterile barrier systems; seal strength, leak tests, aging. | Y | P-PKG-11607; Test methods F88/F1929/F2096; Aging per ASTM F1980 | Seal strength reports, dye penetration, bubble leak, aging protocols/results | PC | Accelerated aging assumes 10°C rule without justification; missing real-time aging initiation. | M | US, EU | N |  | Packaging Engineer | 2025-11-20 | Document Q10 method justification; start real-time aging lot; revise protocol. | Updated protocol; RT aging lot initiated; report addendum approved. | CAPA-2025-031 | Doc review | All sterile SKUs
7.6 (Monitoring & measuring) | Calibrated equipment; traceability to SI; records maintained. | Y | P-760 Calibration; Asset register | Calibration certs, out-of-tolerance assessments, sticker status | C |  | M | US, EU, IR | N |  | Metrology Tech | — | — | — |  | Observation + record sampling | √N+1 per equipment family
8.2.1 (Feedback/Complaints) | Establish documented feedback process including complaint handling. | Y | P-821 Feedback & Complaints; P-853 Corrective Action | Complaint log, investigation forms, trend analysis | PC | Trending not separated by market/device family; no trigger thresholds defined. | M | US, EU, IR | N |  | Post-Market Lead | 2025-10-30 | Define trigger thresholds; split trends by market/device family; link to CAPA triggers. | SOP updated; dashboards show segmented trends; triggers tested. | CAPA-2025-034 | Doc review + data check | Last 12 months of complaints
8.2.2 (Internal audit) | Planned internal audits; criteria/scope/frequency; auditors independent. | Y | P-822 Internal Audit; Annual audit program | Plans, checklists, reports, follow-up/closure | PC | Audit program not risk-based; supplier audits not integrated in the plan. | M | US, EU, IR | Y | Internal | Lead Auditor | 2025-09-30 | Adopt risk-based scheduling; include supplier audits; train auditors on QMSR focus. | New program approved; first risk-based audit executed. | CAPA-2025-037 | Doc review + interview | At least 1 process from each clause group
8.5.2 (Corrective action) | Eliminate causes of nonconformities; documented method; effectiveness check. | Y | P-852 CAPA; Root cause templates (5-Why/Fishbone) | CAPA log, root cause analyses, effectiveness verification | NC | Effectiveness checks missing in 3/7 CAPAs; closure done without sustained evidence. | H | US, EU, IR | Y | Internal | CAPA Coordinator | 2025-10-10 | Define measurable EIs; implement 30/60/90-day checks; retrain owners. | All open CAPAs show EIs achieved with records; SOP updated. | CAPA-2025-041 | Record sampling | All open CAPAs + 3 closed
Labeling (ISO 15223-1:2021/Amd 1:2025) | Use correct symbols and IFU information; update Authorized Representative symbol per 2025 amendment. | Y | P-LBL-15223; IFU templates | Artwork files, symbol legend, label approvals | PC | Old EC-REP symbol still present on two SKUs; legend not updated. | M | EU | N |  | Regulatory Affairs | 2025-09-25 | Update symbol to 2025 amendment; revise IFU/label legends; implement artwork checklist. | Revised artworks approved; checklist in use; print proofs verified. | CAPA-2025-045 | Doc review + sample inspection | All SKUs for EU market