## VMP Outline
Section | Details
Scope & Objectives | Products/sites covered; markets (US/EU/IR); validation policy; SAL target for sterile (1E-6)
Validation Strategy | Risk-based; worst-case approach; DOE where applicable; sampling plans & AQL logic
Roles & Responsibilities | QA/RA, Production Eng., Sterilization/Packaging Eng., Metrology, MR
Process List & Criticality | Register of special processes; CTQs; linkage to risk file
Qualification & Validation | URS->DQ->IQ->OQ->PQ flow; acceptance criteria; requalification frequency
Change Control & Revalidation | Triggers (equipment/material/param/site SW); impact assessment; re-V&V
Documents & Records | Protocols, reports, raw data, deviations, approvals, retention
Monitoring & Trending | SPC plan, control charts, alarms, reaction plan

## Special Process Reg.
Process | Standard/Ref | CTQ | Validation | Requal (months) | Owner
Steam Sterilization | ISO 17665 | F0, Temp, Time, Pressure | IQ/OQ/PQ + BI | 12 | Sterilization Eng.
Radiation Sterilization | ISO 11137 | Dose, Dose rate, Mapping | Dose audit, VDmax/Method 1 | 12 | Sterilization Eng.
EO Sterilization | ISO 11135 | EO conc., Humidity, Temp, Time, Aeration | BI + residuals (ISO 10993-7) | 12 | Sterilization Eng.
Sealing (Packaging) | ISO 11607 | Seal temp/pressure/time | IQ/OQ/PQ + F88/F1929/F2096 | 12 | Packaging Eng.
Cleanroom | ISO 14644 | Particle/microbial counts | Classification + PQ | 12 | Facilities
Cleaning | Internal/ISO | Residue limits | Recovery studies + limits | 24 | Process Eng.

## IQ Checklist
Equipment/Line | URS Met? (Y/N) | Docs Verified | Utilities | Safety/Guards | Calibration | Software/Firmware | Status
Heat Sealer #1 | Y | Manuals, drawings, wiring | Power/air verified | OK | Temp/pressure sensors cal | Version recorded | Pass
EO Sterilizer #2 | Y | P&ID, interlocks, alarms | Steam/vacuum/EO supply | Interlocks tested | Sensors/cal gas ok | SCADA version | Pass

## OQ Protocol
Equipment/Process | Parameter | Low | Nominal | High | Replicates | Acceptance
Heat Sealer | Temperature | 150 | 170 | 190 | 10 | All seals ≥ LSL; no dye penetration
Heat Sealer | Pressure | 2 | 2.5 | 3 | 10 | As above
EO Sterilizer | EO conc. (mg/L) | 500 | 650 | 800 | 3 | BI negative; parametrics in spec

## PQ Plan
Process | Product/SKU | Worst-case Config | Sample Size | Tests | Acceptance
Heat Sealer | Tyvek Pouch 305x400 | Max web width, min dwell | 32 | F88 (strength), F1929 (dye) | Mean ≥ X N; 0 critical leaks
EO Sterilization | Catheter A | Max load density | 3 | BI sterility, residuals (10993-7) | All BI negative; residues ≤ limits

## Requalification
Process | Frequency (months) | Trigger-based? | Next Due | Owner
Heat Sealer | 12 | Yes (change/repair/NC) | 2026-09-01 | Packaging Eng.
EO Sterilizer | 12 | Yes | 2026-02-01 | Sterilization Eng.

## Maint & Cal
Asset | PM Freq | Calibration Freq | Spare Parts | Owner
Heat Sealer #1 | Monthly | 6 mo | Heat bar, Teflon, thermocouple | Maintenance
Pressure Gauge G-12 | - | 12 mo | - | Metrology

## Line Clearance
Check | Status | By | Date
Previous materials removed
Specs/WI latest revision at line
Equipment status OK (cal/PM)

## Batch Release
Item | Status | By
DHR complete & signed
All in-process & final tests passed
Deviations closed/approved

## Parametric Release
Method | Critical Params | Evidence | Decision Rule
Steam | Temp/Pressure/Time, F0 | Cycle records + BI (if applicable) | All parametrics within ranges; alarms none
EO | EO conc., humidity, temp, exposure time, aeration | Cycle data + BI + residuals trending | Within spec; BI negative
Radiation | Dose, mapping | Dose map + Dosimeters readout | Dose within spec window

## Packaging Matrix
SKU | Seal Strength (F88) | Leak (F1929/F2096) | Aging (ASTM F1980) | Shipping/Handling
Kit A | Mean ≥ X N; no critical | No dye/bubble | Accelerated + start real-time | ISTA 2A/3A (if applicable)
Catheter A | Mean ≥ X N | No critical | Real-time started | ISTA as needed

## SPC Plan
Process | Chart | USL | Target | LSL | Sampling | Reaction Plan
Sealing Temperature | Xbar-R | 190 | 170 | 150 | 1 per 30 min | Stop & adjust if 2 of 3 near limit
EO Humidity | Individuals | 65 | 55 | 45 | Continuous | Hold lots outside window

## MSA GageRR
Part ID | Operator | Trial 1 | Trial 2 | Trial 3